|  | ANGLO-AUSTRALIAN OBSERVATORY P.O. BOX 296 EPPING N.S.W. 2121 AUSTRALIA Telephone: +61 2 9372 4800 Fax: +61 2 9372 4880 |
| --- | --- |

- FLAMES FIBRE POSITIONER

| Document Number: | INS-TRE-AUS-13721-0075 |
| --- | --- |
| Issue Number: | 3.2 |
| Issue Date: | 21st Feb 2003 |

OzPoz Calibration Software Manual

| Prepared: | Tony Farrell, Scott Smedley | 21st Feb 2003 |  |
| --- | --- | --- | --- |
|  | Name | Date | Signature |
| Approved: | Tony Farrell | 21st Feb 2003 |  |
|  | Name | Date | Signature |
| Released: | Chris Evans | 21st Feb 2003 |  |
|  | Name | Date | Signature |

Page Intentionally Left Blank

- Change Record

| Issue / Revision | Date | Section / Page affected | Reason / Remarks |
| --- | --- | --- | --- |
| 1.0 | 6th July 2001 | All | First Issue |
| 2.0 | 12th April 2002 | 4 | Added mdlpars, loadhd and savehd commands. Change header formatting. Note changes in names of telescope model parameters in FITS headers. Add section of the model being fitted. Tony Farrell. |
| 2.1 | 9th May 2002 | All | Support wavelength modelling, better selection of items to fit. Explain model and fitting a bit better. Offsets now given in arc-seconds argument then radians. Tony Farrell |
| 2.2 | 18th Jun 2002. | 3.4 | Added section summarising parameters. Tony Farrell. |
| 2.2 | 19th Jun 2002. | 4.7, 4.7.4 (added), 4.8, 4.8.1, 4.8.2, 4.8.4 (added), 4.8.5 (added) | Add “show fit” and “show model” commands. Add “set fit n=[no]fit” command. Add some comments etc. on operations and references to other relevant commands. |
| 2.2 | 19th Jun 2002 | 3.4.2, 3.4.4. | Modified for updated telescope model. |
| 2.3 | 27th Jun 2002 | Several | Remove “set linear” command and “create” command. Warning about problems in fitting linear parameters and not removing origin. Add “extract” command. |
| 3.0 | 19th Oct 2002 | 4.7, 4.6.3, 4.12, 4.13, 4.6, 4.6.7, 4.6.8. | Note – I have lost an intermediate update due to file corruption. Add “clear”, “show residuals”, “set fit rod”, “set powell”, “set fitfun” “remmeanoff” commands. Describe powell fit function. New options to fluxOffsets command. |
| 3.1` | 20th Oct 2002 | Added 6 | Describe techniques used during comm4. |
| 3.2 | 21st Feb 2003 | Cover Only | Apply Signatures. |

Page Intentionally Left Blank

- OzPoz Calibration Software Manual

1	References	9

1.1	Applicable Documents	9

1.2	Reference Documents	10

1.3	Abbreviations and Acronyms	10

1.4	Glossary	10

2	Introduction	11

3	Software	11

3.1	How the calibration software is used	11

3.2	Inputs	12

3.2.1	FITS File Format	12

3.2.2	ASCII Offset File Format	13

3.2.3	ASCII Fibre Fluxes File Format	14

3.2.4	Image Positions File Format	15

3.3	Output	15

3.4	The fitted model.	15

3.4.1	Linear Model.	15

3.4.2	Distortion Model.	16

3.4.3	Field Plates	16

3.4.4	Model Parameter Summary	16

3.5	Wavelength	18

3.5.1	How Wavelength is handled.	18

3.5.2	Fitting Wavelength Effects	19

4	User Commands	19

4.1	load	20

4.2	fit	21

4.2.1	Powell Fit	21

4.2.2	SVD Fit	21

4.3	help	21

4.4	offsets	22

4.5	quit	22

4.6	set	22

4.6.1	set offset <button#> <RA> <DEC>	23

4.6.2	set fit <0> <2> <3> <4> <5> <6> …<n>	23

4.6.3	set fit nolnear|[no]origin|[no]rotation|[no]scale|[no]dist|[no]wdist|noall	23

4.6.4	set fit n=0|1|[no]fit	23

4.6.5	set parameters <0> <2> <3> <4> <5> <6>…<n>	24

4.6.6	set wavelen giraffe|uves|facb <wavelength>	24

4.6.7	set fitfun powell|svd	24

4.6.8	set powell tolerance|maxits <value>	24

4.7	show	24

4.7.1	show fpcalData	24

4.7.2	show model	25

4.7.3	show offsets	25

4.7.4	show pars	25

4.7.5	show fit	25

4.7.6	show residuals	25

4.8	mdlpars	25

4.8.1	mdlpars loadsds	26

4.8.2	mdlpars loadvlt	26

4.8.3	mdlpars show	26

4.8.4	mdlpars rotate	26

4.8.5	mdlpars scale	26

4.8.6	mdlpars savesds	26

4.8.7	mdlpars savevlt	26

4.8.8	mdlpars set	26

4.9	loadhd	26

4.10	savehd	27

4.11	extract	27

4.12	clear	27

4.13	remmeanoff	28

5	Utilities	28

5.1	genOffsets	28

5.2	fluxOffsets	29

5.3	img2flux	29

5.4	region	30

5.5	residuals.tcl	31

5.5.1	commands	32

5.5.2	options	33

5.6	cal.sh	33

6	Comm 4 Approach.	35

6.1	Input Data	35

6.1.1	Extracting information from Headers.	35

6.1.2	First and last flux files.	35

6.2	Determining best offset from fluxes.	35

6.3	Basic Analysis	36

6.4	Distortion analysis.	37

6.4.1	Wavelength distortion terms.	37

6.5	Putting new values into the Robot.	37

6.5.1	Raw fit results into Robot.	37

6.5.2	Hand edited results.	37

6.5.3	Passing the new model to configure.	38

6.6	Software location and building.	38

Page Intentionally Left Blank

- Fibre Positioner

- OzPoz Calibration Software Manual

# References

## Applicable Documents

The following documents, of the exact issue shown, form a part of this document to the extent specify herein.

- VLT-SPE-ESO-10000-0011, 2.0 30/09/1992 -- VLT Software Requirements Specification

- VLT-SPE-ESO-17212-0001, 2.0 12/04/1995 -- VLT Software - VLT Instrumentation Software Specification

- VLT-SPE-ESO-17240-0385, 2.1 15/07/1996 -- VLT Software - INS Common Software Specification

- VLT-PRO-ESO-10000-0228, 1.0 10/03/1993 -- VLT Software - Programming Standards

- VLT-MAN-SBI-17210-0001, 3.4 05/05/97 -- VLT Software - LCU Common Software, User Manual

- VLT-MAN-ESO-17210-0690, 3.3 30/10/98 -- VLT Software – Graphical User Interface, User Manual

- VLT-MAN-ESO-17210-0667, 1.0 03/12/96 -- VLT Software – Guidelines for the Development of VLT Application Software

- VLT-MAN-ESO-17210-0707, 1.5 30/04/97 -- VLT Software – CCS On Line Database Loader, User Manual.

- VLT-MAN-ESO-17210-0667, 1.0 03/12/96 -- VLT Software – Guidelines for the development of BLT Application Software.

- VLT-PLA-ESO-00000-0006, 2.0 21/05/92 – VLT Program – Software Management Plan.

- INS-SPE-ESO-13720-1600, 1.0 06/05/1998 -- VLT Instrumentation Plan - FLAMES Fibre Positioner, Technical Specification

- INS-SPE-AUS-13721-0031, 2.0 29/09/1999 -- Fibre Positioner, Software Requirements and Software Functional Analysis

- INS-SPE-AUS-13721-0042, 1.0 29/09/1999 -- Fibre Positioner, Instrument Control Software Design Description

- VLT-SPE-ESO-13730-1816, 1.0 12/04/1999 – Flames Observation Software Design Description

- INS-PLA-AUS-13721-0035, 2.0, 29/09/1999 – Flames Positioner Calibration and Commissioning Plan.

- VTL-PLA-ESO-13700-1877, 1.0, xx/xx/1999 – Flames Operation and Calibration Plan. (in preparation)

## Reference Documents

The following documents are referenced in this document:

- INS-INS-AUS-13721-0017, 4 17/11/1997 -- Flames Fibre Positioner, Glossary of Terms

- INS-SPE-AUS-13721-0071, 1.1 08/05/2001 – Astrometric Model and Associated Software.

## Abbreviations and Acronyms

The following abbreviations and acronyms are referenced in this document:

| AAO | Anglo-Australian Observatory |
| --- | --- |
| Dec | Declination |
| FLAMES | The Whole Fibre Facility |
| FP | Fibre Positioner |
| GUI | Graphical User Interface |
| RA | Right Ascension |
| Tcl | Tool Command Language |
| Tcl/Tk | Tcl Windowing Toolkit. |

## Glossary

A glossary of terms specific to the Fibre Positioner is given in the separate document Flames Fibre Positioner, Glossary of Terms [17].

# Introduction

This document describes how to use the OzPoz Astrometric Calibration Software. This software is utilised during commissioning to obtain an accurate Astrometric Model for the OzPoz instrument. The model is subsequently used for mapping the field plate position of a source object (in the field of view) with the object's mean RA/Dec coordinate and vice-versa.

To calibrate the OzPoz instrument, a series of commissioning tests are performed with different instruments (described in detail in [18]) and the results passed to this software. The results are analysed and an Astrometric Model built. The rest of this document describes how to use this software to build the model.

# Software

## How the calibration software is used

Below is a brief description of how to use the calibration software to create an accurate Astrometric model.

- Position a number of fibres on the field plate where stars of accurately known positions are predicted to be.

- Save the state of the telescope and fibres in a FITS file.

- For each fibre, manually adjust the position of the telescope until the light intensity coming from its star is at a maximum. Then, record the fibre number and offset RA/Dec (radians) of the telescope from its initial position. This can be saved to an ASCII Offset File (see § 3.2.2) or manually entered into the fpcal program by the set buttonoffset command (see § 4.6.1).

- Fire up the fpcal program if you haven't already done so.

- Read in the FITS file of observations that you saved in step 2:

load <FITSfile>

where <FITSfile> is the name of the file you saved the observations as.

- If you chose to save the telescope offsets in a file in step 3, load the offsets with the command:

offsets <filename>

where <filename> is the name of the file you saved them in.

- Confirm that the telescope offsets are correct. View the current values with:

show offsets

and make any alterations with the set buttonoffset command described in § 4.6.1.

- Set the parameters you wish to fit

set fit <parameters>

- Create an Astrometric model based on the initial model and the offsets:

fit

Upon successful completion, an SDS file containing the astrometric model will be saved. You will be told the name of this file.

- Check that the newly created astrometric model produces reasonable results. You can look at a graphical representation of the telescope, fibre buttons, offsets and residuals with the residuals.tcl utility. (see § 5.5) Note that this is a separate program and NOT a command of fpcal.

residuals.tcl residuals.out

The residuals.out file is automatically generated after every successful fit command in fpcal. i.e. it should have been generated for you in step 9.

## Inputs

To calculate the parameters for the Astrometric Model, the calibration software needs to know the RA/Dec of each fibre in the field (positioned on an astrometric star) and how much the telescope was aligned to centre the relevant star on each fibre. (We refer to the direction and magnitude of each alignment as an offset.)

The software also requires a number of context parameters such as: field centre, atmospheric conditions, date of observation, etc. Most of this data comes from the FITS file output by the instrument making the observations. The FITS file format and necessary fields are described in detail in § 3.2.1.

Currently, the calibration software makes use of 2 different input files:

- a FITS File containing a Binary Table, and

- an ASCII offset file

Two utilities exist for generating ASCII Offset files. Depending on the output data of your instrument, they may be useful to you. Refer to § 5.1 and § 5.2 for a description of their purpose and operating instructions.

### FITS File Format

The calibration software expects a FITS file with a binary table that has an extension name of "OzPoz_table". The binary table must contain the following header fields:

- CENRA		Field Centre Right Ascension (DDMMSS.SS radians)

- CENDEC		Field Centre Declination (DDMMSS.SS radians)

- ATMTEMP	Atmospheric Temperature (Kelvins)

- ATMPRES	Atmospheric Pressure (millibars)

- ATMRHM	Atmospheric Relative Humidity (percentage)

- GIRAWLEN	Giraffe spectrograph wavelength (nanometres)

- UVESWLEN	UVES spectrograph wavelength (nanometres)

- FACBWLEN	FACB  wavelength (nanometres)

- ACTMJD		Observation Date (Mean Julian Date)

- POSANGLE	Position Angle of Y Axis (radians)

- CENEQNX	Equinox of Observation Date

- PM_RA		Proper Motion Right Ascension (DDMMSS.SS radians)

- PM_DEC		Proper Motion Declination (DDMMSS.SS radians)

- TELPAR1		First parameter of model in observing instrument.

- TELPAR2		Second parameter of model in observing instrument.

- TELPAR3		Third parameter of model in observing instrument.

- TELPAR4		Fourth parameter of model in observing instrument.

- TELPAR5		Fifth parameter of model in observing instrument.

- TELPAR6		Sixth parameter of model in observing instrument.

- TELPARn		nth parameter of model in observing instrument.

The binary table must contain the following fields:

- BUTTON		Fibre button number.

- RA		Mean RA of astrometric star (DDMMSS.SS radians).

- DEC		Mean declination of astrometric star (DDMMSS.SS radians).

### ASCII Offset File Format

The ASCII Offset file simply contains a list of fibre numbers with an associated telescope offset RA/Dec (in arc-seconds (note – this has changed, it was radians). The offset represents the amount and direction of deviation between the position a star in the field of view was thought to be located at, and the actual position it was found at.

It is implemented as a (column-space-separated) table with 3 columns:

- column 1 - fibre number (integer)

- column 2 - offset RA in arc-seconds (double)

- column 3 - offset Dec in arc-seconds (double)

Note that a dummy offset file can be generated for you with the genOffsets utility (described in § 5.1) and that the fluxOffsets utility will generate a valid ASCII Offset File from a set of ASCII Fibre Fluxes files. (see § 5.2)

The offsets file can also have entries to set the wavelength for the given fibre.  These lines cause entries for offsets to be associated with the new wavelength rather then the old.  See the section on wavelengths below for full details.  The software maintains one wavelength for each fibre type.  The fibre type is determined automatically from the fibre number and is one of UVES, GIRAFFEE or FACB.  The relevant lines in the offset files would be as follows

GIRAFFE_WLEN <wavelength>

UVES_WLEN <wavelength>

FACB_WLEN <wavelength>

Where <wavelength> is given in nanometers.  If not given, the wavelengths associated with an offset are the current wavelength for the specified time, as read from the binary table header or set with the “set wavelen” command.

### ASCII Fibre Fluxes File Format

The first line of an ASCII Fibre Fluxes File contains the telescope RA/Dec offset at which an instrument has made its observations. Both values are in radians and represented as the C type, double.

The remainder of the file is a table with two columns:

- column 1 - Button number of fibre (integer)

- column 2 - flux in fibre in photons/sec (double)

Some versions of this file may contain a 3rd and 4th column denoting the flux in different wavelengths. See § 5.2 for a description of how this information can be used.

### Image Positions File Format

This file is used to specify the position of fibre regions and background regions in a UVES slit image. See img2flux (§ 5.3) and region (§ 5.4) for a description.

The file comprises a table with 5 columns:

- column 1 – Fibre number (-1, -2 for background region) (integer)

- column 2 – x coordinate of top-left corner of bounding box (integer)

- column 3 – y coordinate of top-left corner of bounding box (integer)

- column 4 – x coordinate of bottom-right corner of bounding box (integer)

- column 5 – y coordinate of bottom-right corner of bounding box (integer)

## Output

The 2 significant outputs of the calibration software are the Astrometric Model and the residuals.

The Astrometric Model is stored in SDS format by means of a call to fpmdlWriteModel. This SDS file can subsequently be loaded into the robot software using the robot software command fpcolMdlLoad.  Note that the current model can be extracted from the robot software using fpcolMdlExtract.  Note that the “mdlpars” command has sub-commands which can help here.

The residuals output file can be loaded into a separate program, residuals.tcl, to view the calculated residuals and offsets graphically. It is useful to diagnose errors in the data & verify that the Astrometric Model being generated is correct. See § 5.5 on the residuals.tcl utility for further details.

## The fitted model.

The software fits a linear model and a distortion model.  These are parameterised and it is these parameters (known as the telescope model parameters) which the fitting routines are trying to optimise.  Currently the sofware allows for 20 such parameters, of which 14 are currently used.  The remainer allow for expansion if required.

### Linear Model.

The linear model parameters are the first six telescope model parameters.  They are used as follows

x’  =  a + bx + cy

y’  =  d + ex + fy

Thus  pars[0] = a,  pars[1] = b etc.

The next two parameters are used to allow independent fitting of rotation and scale.  These are not normally used by the real model, just in the fitting software.  When the model is written to a file, pars[6] (rotation in degrees) and pars[7] (scale) are first applied to the above linear model and then rest to 0 and 1 respectively.

WARNING:It makes little sense to independently fit parameters b, c e and f.  The software deliberately does not make this easy to do and will warn if you attempt to do so.  The preferred approach is to fit the origin first, and then the scale and rotation.

Also note that the distortion model below also includes a scale component so fitting both scale and the distortion model should be avoided.

### Distortion Model.

The next 6 parameters are used for the distortion model is as follows.

R = a * th + b * th3 + c *th5 + d * th/l + e * (th/l)2

Thus pars[8] = a, pars[9] = b etc.  Where “th” represents the off axis angle and l the observing wavelength.

Note that it makes little sense to try to model all of these in one operation – the three that are dependent on wavelength will impact the three not dependent on wavelength.  Also, trying to fit a such a large number of parameters requires a significant data set

### Field Plates

It should be noted that the model does not mention the field plate.  Currently a different set of model parameters is applied to each field plate, allowing for differences between the plates (for example, the plate rotation zero point).  During commission it is likely that as a first pass, the parameters of one plate will be determined and then loaded for the other plate.  This can be done by copying the resulting file (OzModel<n>.sds, n = plate number) to the file for the other plate before loading it into the robot software.

### Model Parameter Summary

The following table summaries the various parameters, with units and descriptions.

| Parameter | Units | Description |
| --- | --- | --- |
| 0 | Microns | Parameter a of the linear model. Offset of the plate centre in X. Not expected to be used (should remain zero) as field acquisition will remove this error automatically. |
| 1 | None | Parameter b of the linear model. Significant effects incorporated are the scale and the cosine of the rotation. If allowed when fitted, then will also contain components modelling squash and shear. |
| 2 | None | Parameter c of the linear model. Significant effects incorporated are scale and the sine of the rotation. If allowed when fitted, then will also contain modelling squash and shear. |
| 3 | Microns | Parameter d of the linear model. Offset of the plate centre in Y. Not expected to be used (should remain zero) as field acquisition will remove this error automatically. |
| 4 | None | Parameter e of the linear model. Significant effects incorporated are scale and the negative of the sine of the rotation. If allowed when fitted, then will also contain components modelling squash and shear. |
| 5 | None | Parameter f of the linear model. Significant effects incorporated are scale and the cosine of the rotation. If allowed when fitted, then will also contain components modelling squash and shear. |
| 6 | Degrees | This parameter is not part of the final model. It is required to allow the modelling of rotation independent of scale, squash and shear. If you want to fit only rotation, then fit this parameter. When the model is written out, this value will be applied to the linear model and then set to zero. |
| 7 | None | This parameter is not part of the final model. It is required to allow the modelling of scale independent of rotation, squash and shear. If you want to fit only scale, then fit this parameter. When the model is written out, this value will be applied to the linear model and then set to 1. |
| 8 | mm/degree | Parameter a of the distortion model, also known as the first non-wavelength dependent distortion model parameter. |
| 9 | mm/degree3 | Parameter b of the distortion model, also known as the second non-wavelength dependent distortion model parameter. |
| 10 | mm/degree5 | Parameter c of the distortion model, also known as the third non-wavelength dependent distortion model parameter. |
| 11 | mm.nm/degree | Parameter d of the distortion model, also known as the first wavelength dependent distortion model parameter. |
| 12 | mm.(nm/deg)2 | Parameter e of the distortion model, also known as the second wavelength dependent distortion model parameter. |
| 13 … 19 | N/A | Reserved for future expansion. |

## Wavelength

Wavelength is an issue of significant complexity for this software.  The observation wavelength is used in two ways by the telescope modelling software.  First, the wavelength is used in the refraction and atmospheric dispersion calculations.  These parts of the calculations are done using well-known formulae and parameters.  The wavelength is also used in the distortion model adjustments, as seen above.  It is here that there are wavelength dependent parameters that must be fitted by this software.

The instrument supports different wavelengths for different fibre types.  Currently there are three relevant fibre types, fibres feed to the Giraffe Spectrograph (Medusa, IFU’s, ARGUS Sky), fibres feed to the UVES spectrograph and the FACB fibres used for guiding.

### How Wavelength is handled.

Each offset loaded by the program has a wavelength associated with it.  There can be a maximum of ten different wavelengths known to the software at any given time.  (If you exceed this you must stop and restart the fpcal program).  Each offset also has an object associated with it (details read from the file header).  So each offset known by the program is a doublet of button (and hence object) and wavelength.

There is a current wavelength for a given fibre type.  The initial wavelength for each type comes from the header file.  It can be changed using “set wavelength” or in the offset file.  When an offset for a button is loaded, it is associated with the currently wavelength for that fibre type – creating a doublet as mentioned above.  To change an offset for a given wavelength and button combination, the wavelength must be the current wavelength for the given fibre type.

When fitting, the object of the exercise is to generate a set of model parameters that best fits the data.  For each of the known wavelengths, the software maintains an initialised telescope model.  When calculating with these models, the same set of model parameters is used for all wavelengths.  This allows the fitting algorithm to vary model parameters such as to optimise a fit across multiple wavelengths.

### Fitting Wavelength Effects

First note that the wavelength effects are relatively small, and it may be hard to optimise the parameter much (default values were chosen for analysis of the optical design).

To fit the wavelength parameters, you will need data for different wavelengths – at least three and probably more to model the higher order terms.  You should configure and observe the same field at each wavelength (note, you must re-configure to take account of atmospheric effects).  For each wavelength, determine the set of offsets required to maximise the flux in each fibre.

You should now have an offset file for each wavelength.  Ensure each file is prefixed by lines of the form (“GIRAFFE_WLEN <wavelength>”, “UVES_WLEN <wavelength>” or “FACB_WLEN <wavelength>”).  Then concatenate these files into one file.

You can now run fpcal.  The FITS file from which the binary table is extracted can be that associated with any one of the observations (they should all be specifying the same information, except for wavelength and observation time).

You can now load your concatenated offset file – giving you multiple offsets for each fibre, each associated with a different fibre.  Set the parameter to be fitted and run the fit.

# User Commands

fpcal is used to generate the Astrometric Model and allows the user to interactively set and vary any of the parameters.

The available commands are:

- load	Load a FITS file.

- fit		Fit current parameters and offsets. Creates an Astrometric Model.

- help	Simple help system.

- offsets	Load offsets from a file.

- quit/exit	Exit the program.

- set		Alter some of the loaded parameters.

- show	Show the current value of specified parameters.

- mdlpars	Manipulate the telescope model parameters.

- savehd	Save the header details loaded by “load” or “loadhdr” in a text file.

- loadhd	Load header details from a text file of the type written by “savehd”.

- extract	Extract data for an external fit.

- clear	Clear’s button offsets

- remmeanoff Remove the mean offset

You can alter the values of most of the data loaded from the FITS and offsets files, however you must be aware of dependencies. Some commands require fpcal to be in a certain state before it can continue. fpcal will notify you if you try to violate one of these dependencies. For example, a FITS file and offset file must be loaded before you can create a model.  Additionally, each of the standard commands will indicate the commands that normally follow, with the first command in such output being the normal command which follows.

Dependencies are listed with each command below.

## load

usage: load <FITSFile>

dependencies: none

This command will read in all the relevant data contained in the specified FITSFile. You can view the values of this data with the show fpcaldata command.

Specifically, it loads the following parameters:

- atmTemp		Atmospheric temperature at time of observation.

- atmPressure	Atmospheric pressure at time of observation.

- atmHumidity	Atmospheric humidity at time of observation.

- wavelength	Wavelength observation is interested in.  You can fit for multiple wavelengths if required and different wavelengths may be associated with each offset.

- equinox		Calendar system of observation date.

- posAngle		Position angle of observing telescope.

- fieldPlate		Field plate number.

- observationDate	Date of observation.

- nFibres		Number of fibre buttons on field plate.

- fieldCentre	RA/DEC of field centre of telescope.

- apparent		Apparent RA/Dec of fibres (converted from mean RA/Dec).

- parameters	The telescope model parameters as used for the observations.

- fit			Parameters to fit when creating model. [default: all]

You can alter some of these values. Refer to the set command (§ 4.6).

## fit

usage: fit

dependencies: offsets, load/loadhd “set fitfun” and “set fit”

This command will create an Astrometric model based on the, the current state of the data loaded from the FITS file and the offsets.

Upon successful completion, fit will generate 2 output files:

| - OzPozModelN.sds | SDS file containing fitted Astrometric Model. N is the field plate number from the header of the binary table in the FITS file. |
| --- | --- |
| - residuals.out | File containing calculated residuals suitable for input to the residuals.tcl utility. |

The default fit method is the SVD Fit.

### Powell Fit

This is a Powell minimization using an algorithm from Numerical Recipes.  The Powell fit has two parameters that can be changed, the tolerance on the final error and the number of iterations.  See the set command for details.

Currently – it is not clear if the powell fit is any better then the much quicker SVD fit.  Additions, you get a lot of warnings during execution that should be ignored.  The Powell fit does not currently give variances on the results.

### SVD Fit

This is the SLA library SVD fitting algorithm. The fit is done using “Singular value decomposition”, using the sla library routines SLA_SVD(), SLA_SVDSOL() and SLA_SVDCOV().  Please see http://www.starlink.rl.ac.uk/star/docs/sun67.htx/sun67.html#xref, particularly the section on :”Numerical Methods:

It should be noted that the fit algorithm automatically restrains spuriously large coefficients.  This may happen for example if there is insufficient data to correctly fit those parameters.

## help

This command will list all the available commands in the fpcal program. You can obtain further help by typing:

<command> help

where <command> is one of the available commands.

## offsets

usage: offsets <file>

dependencies: none

This command will load the offsets specified in <file> into fpcal ready for use by the 'fit' command. Note that once the offsets have been loaded they can be viewed with:

show offsets

and altered with:

set offset <button#> <RA> <DEC>

## quit

usage: quit

dependencies: none

Exit the fpcal program.

## set

usage:

set offset <button#> <RA> <DEC>

set fit <0> <2> <3> <4> <5> <6>… <n>

set fit nolnear|[no]origin|[no]rotation|[no]scale|[no]dist|[no]wdist|noall

set fit n=0|1|[no]fit

set parameters <0> <2> <3> <4> <5> <6> …<n>

set wavelen giraffe|uves|facb <wavelength>

set fitfun powell|svd

set powell tolerance|maxits <value>

dependencies: set fit and set parameters both depend on load or loadhd.

This command allows the user to override the fit and parameters values stored in the fpcaldata structure as well as the button offsets.

### set offset <button#> <RA> <DEC>

Set the offset RA to <RA> and offset Dec to <DEC> for fibre button number <button#> (in arc-seconds).

You can view the button offsets with show offsets.

### set fit <0> <2> <3> <4> <5> <6> …<n>

Specify which model parameters are to be used in the fit. For each of the parameters, a zero means the parameter will be fitted, a 1 means it is fixed. The default is not to fit any parameters – you must explicitly say which parameters you wish to fit.  It is normally easier to specify the items to fit using the keyword version of this command.  See below.

The state of the fit parameters can be viewed in the fpcaldata structure with show fpcaldata. Look for the “fit:” entry.

### set fit nolnear|[no]origin|[no]rotation|[no]scale|[no]dist|[no]wdist|noall

Short hand for setting/clearing particular fit items

[no]linear	Don’t fit the linear model parameters (index 0 through 5).  There is no positive version of this command as it makes little sense to fit the linear model parameters in most cases.  You enable fitting them using other forms of the fit command.

rod	Fit the rotation, origin and distortion.

[no]origin	Disable/Enable fitting the origin.

[no]rotation	Disable/Enable fitting for rotation.

[no]scale	Disable/Enable fitting for scale.

[no]dist	Disable/Enable fitting for the non-wavelength dependent distortion coefficients.

[no]wdist	Disable/Enable fitting for the wavelength dependent distortion coefficients.

noall	Disable fitting for all coefficients.  Useful for returning to the default state.

### set fit n=0|1|[no]fit

Indicate if a particular model parameter is to be fited, via index.  “n” is the index of the parameter to fit.  The value is either 0 – indicating the parameter is to be fited, 1 indicating is it not, or the text equivalents fit/nofit.

### set parameters <0> <2> <3> <4> <5> <6>…<n>

Override the parameters in the fpcaldata structure as loaded from the FITS file.  Also see the “mdlpars set” for other ways to change these values.

The state of parameters can be viewed in the fpcaldata structure with show fpcaldata. Look for the “parameters:” entry.

### set wavelen giraffe|uves|facb <wavelength>

Set the wavelength to be assumed for future offsets for fibres associated with the given instrument (Giraffe, UVES or FACB’s).  Note that this may be overridden in an offsets file.  Does not impact any already loaded offsets.

### set fitfun powell|svd

Select the type of fitting function to use.  Either a Powell minimization or singular value decomposition (currently the default).

### set powell tolerance|maxits <value>

Set the powell fit tolerance (microns, default 0.01) or maximum iterations (default 200).

## show

usage:	show <var>

where <var> is one of:

- fpcalData

- model

- offsets

- pars

- fit

- residuals

- all

dependencies: none

### show fpcalData

This command will display all the data in the fpcaldata structure. If the load command has not been successfully called, it will simply display that the structure has not been initialised.

The data shown here includes the data for “show pars” and “show fit”.

### show model

This command will display whether or not an initial model has been created.  It will also show the current wavelengths know by the system

### show offsets

This command will list the offset RA/Dec for each button.

If no offsets have been loaded, or none specified with the set command, it will indicate that the offsets have not been initialised.

### show pars

Show the current values of the telescope model parameters.  Includes the value, a simple description and units.  See also the “mdlpars show” command, which decomposes some components.

### show fit

Indicate if parameters are to be fited or not.

### show residuals

Output the residuals generated in the last fit operation.

## mdlpars

usage:	mdlpars loadsds <plate #>

mdlpars loadvlt <plate #>

mdlpars show

mdlpars rotate <degrees>

mdlpars scale <factor>

mdlpars savesds

mdlpars savevlt

mdlpars set <par> <value>

where 	<plate #> is a FP plate number (1 or 2):

<degrees> is an angle in degrees

<factor> is a real number

<par> is a telescope model parameter index from 0 to n

<value> is a real number.

dependencies: For all except the loadsds/loadvlt versions - load or loadhd or “mdlpars loadsds” or “mdlpars loadvlt”

This command is used to manually change the telescope model parameters.

### mdlpars loadsds

The “loadsds” command will load the model parameters for the specified plate from OzPozModelN.sds.

### mdlpars loadvlt

The “loadvlt” command can only be used on the “wflames” instrument control workstation.  It first causes the FLAMES FP Software to write the file OzPozModelN.sds.  It then loads this file into the fpcal program.

### mdlpars show

The “show” command will show the existing values and some decomposition of the model parameters.

### mdlpars rotate

The “rotate” command applies the specified rotation change to the telescope model.

### mdlpars scale

The “scale” command applied the specified scale change to the telescope model.

### mdlpars savesds

The “savesds” commands saves the current model to OzPozModelN.sds.

### mdlpars savevlt

“savevlt” can only be used on the wflames instrument control workstation – where after saving the file OzPozModelN.sds, it causes its contents to be loaded into the FLAMES FP software.

### mdlpars set

The “set” command allow an individual telescope model parameter value to bet set.

## loadhd

usage: loadhd <TextFile>

dependencies: none

This command will read in all the relevant data contained in the specified TextFile.  This is exactly the same information as is read by “load” –see section 4.1.  The format of the file is as is written by the “savehd” command.  The file “example_header.txt” is provided with the software as an example.  The idea is that you can edit this file and provide values with needed in fits file, if required (say for calibrating with the FACB or UVES slit view).

Note that the file format is very fixed – you must stick to that format, just changing the numbers or adding/subtracting objects.

## savehd

usage: savehd <TextFile>

dependencies: load or loadhd

This command will save the details described in the load command – see section 4.1, into a text file.  This can be used just to save this information or to allow it to be edited.  It can be loaded back in use the “loadhd” command above.

## extract

usage: extract <TextFile>

dependencies: load or loadhd and offsets/set offset

This command will write to the specified file a table consisting of the following columns

- The button number of an offset

- The wavelength of the offset (nanometers)

- The Tangent plane “xi” coordinate of the object.(degrees)

- The Tangent plane “eta” coordinate of the object. (degrees)

- The RA offset, radians.

- The Dec offset, radians.

This command is useful for extract the data required to run an external fit program.

## clear

usage: clear <buttonNo> [<index>]

clear resabove <microns>

The first form of this command will clear the particular button’s offset for the current wavelength index for the fibre type.  The optional index argument is a wavelength index to clear the offset for.

The second form of this command will clear all offsets with a residual above the specified value in microns.

## remmeanoff

Removes the mean of the offsets from all the offsets.

# Utilities

## genOffsets

usage: genOffsets [-options] <FITSfile>

available options are:

-h		help

-n <noise>	number of arc seconds to vary the offset by [0.0]

-o <file>	output filename [-]

-p		prompt for parameters to arbitrary model

<FITSfile> is a FITS file containing a binary table with fibre button numbers in it.

The genOffsets utility will generate a dummy/test ASCII Offset file from the specified FITS file. This is useful for testing and ensuring the integrity of the Astrometric Model output by the calibration software.

This utility outputs an offset file based upon the fibre button positions in the specified FITS file, and an arbitrary model. In effect, the offset for each fibre button is the difference between the arbitrary model and the model reconstructed from the PARS parameters in the FITS file.

The -n command-line argument specifies the maximum amount of random noise to add into system. This takes the form of random additions/subtractions in the range ±<noise> arcseconds from the offset location. The greater the noise, the greater the difficulty fpcal will have in creating an accurate Astrometric Model.

Given only the FITS file and offset file generated with little or no noise as input, the calibration software should, ideally, determine an Astrometric Model very close to the arbitrary model chosen to generate the offset file.

To generate offsets for testing wavelength modeling, you will need to run this multiple times, once for each wavelength (must be changed in the header - a text header works better).  Then concatenate the files into one offset file.

## fluxOffsets

usage: fluxOffsets [-option] <file1> <file2> ... <fileN>

options:

-rgb			Input files containing fluxes in RGB wavelengths.

-verbose		Extra logging information.

-o <ouputFile>	Specify location of ASCII offset file.

-gaussfit		Use gaussian fitting routine to find centroid.

-gaussfit2		Use second gaussian fitting routine to find centroid.

-kgfit			Use KG’s gaussian fitting routine.

-gdhs			Input files are in the GIRAFFE DHS format.

-nosubmin		Disable subtraction of minimum values.

-cut <val>		Cut out all offsets with a poor signal to noise ration then the specified value (taken as the ration of the max value to the min value in a grid.  Only valid if –nosubmin is NOT specified

The fluxOffsets utility will generate an ASCII Offset file (see § 3.2.2) from a set of ASCII Fibre Fluxes files (see § 3.2.3). ASCII Fibre Fluxes files are usually generated from UVES slit images with the img2flux utility (see § 5.3).

The utility scans all of the specified Fibre Flux files and determines the offset of the telescope when the measured flux/intensity is at a maximum for each fibre button.

If the -rgb argument is specified, the program will assume that each Fibre Fluxes file contains a flux value for red, central and blue wavelengths.

Only the central wavelength flux values are used in determining the offset RA/Dec for each fibre. However, the program will warn you if the red and/or blue wavelength flux values would give a significantly different offset value. Currently, any deviation of 1 arcsecond or greater  presents a warning.

## img2flux

usage: img2flux [-options] <FITSFile1> … <FITSFileN>

options:

-h 		help

-i <file>	file containing image positions [img.positions]

-p <prefix>	output filename prefix [f]

-p <postfix>	output filename postfix [none]

-v 		verbose output

This program accepts as input a set of FITS files, each containing a UVES slit image (similar to above), and outputs a set of Fibre Flux files (see § 3.2.3). The Image Position file (see § 3.2.4) specified by –i specifies the region each fibre encompasses within the slit image. The program determines the light intensity/flux within each region for a fibre and outputs this information in the format of a Fibre Flux file for each image.

The telescope offset for each image must be stored in a COMMENT record in the header of the associated FITS file.

## region

usage: region [-options] <FITSfile>

options:

-h 		help

-i <file>	file containing image positions [img.positions]

-o <file>	output FITS file [region.fits]

-v 		verbose

This utility is used when wanting to confirm the accuracy of the fibre regions specified in an Image Positions file (see § 3.2.4).

This utility draws the bounding boxes specified in an Image Positions file (-i command-line argument) on a copy of the input FITS file. The input FITS file must contain a UVES slit image.

The resultant output FITS file contains a modified slit image similar to above. Note the presence of 2 large regions above and below the fibre regions – these are the regions used to calculate the background noise in the img2flux (§ 5.3) utility.

## residuals.tcl

usage: residuals.tcl <residuals.out>

This utility will display a graphical representation of the field plate as modelled by the Astrometric Model output by fpcal. It allows the user to visualise the placement of each fibre button on the field plate along with the associated offset and residual (magnitude and direction).

The sole input to this utility is the residuals.out file that is output as the result of a successful fit command in fpcal. The format and contents of this file should not concern the general user and as such no description is contained here. Suffice to say, this file contains all of the information necessary to plot the fibres, offsets and residuals for the Astrometric model in the residuals.tcl utility.

### commands

At the bottom of the residuals.tcl display, there are 3 fields that allow the user to adjust scale/magnification parameters. They are:

| - Residual Magnification | Amount to exaggerate the magnitude of the residual error. [default: 10000] |
| --- | --- |
| - Offset Magnification | Amount to exaggerate the magnitude of the offset error. [default: 100] |
| - Scale | microns/pixel [default: 5000] |

### options

The main options menu has several choices:

| - Reset Mag/Scale | Adjust the magnification and scale back to their default sizes. |
| --- | --- |
| - Centre | Centre the field plate in the middle of the screen. |
| - AutoCentre | Centre the field plate in the middle of the screen after every magnification or scale adjustment. |
| - Show Fibres | Display each fibre button in blue on the field plate. |
| - Show (Plate) Crosshair | Display a red crosshair centred on the field plate. |
| - Show Plate (Boundary) | Display a red circle denoting the boundary of the field plate. |
| - Show Residuals | Display a black vector denoting the residual for each fibre button. Note that this vector is magnified according to the value of the 'Residual Mag.' Field at the bottom of screen. |
| - Show Offsets | Display a green vector denoting the offset for each fibre button. This vector is magnified according to the value of the 'Offset Mag.' field at the bottom of screen. |

Moving the mouse over a fibre button, offset vector and/or residual vector, will popup a list of information about the objects under the cursor. For fibre buttons, the button number and x/y position (in microns), as an offset from the field centre, is displayed. For offset and residual vectors, the associated fibre button number and x/y position (in microns) as an offset from the fibre button position, is displayed.

Offset information in this list is appended with an 'O', and residual information with an 'R'.

## cal.sh

usage: cal.sh [ options ] <FITSfile>

available options are:

-h help

-n <noise> amount of random noise to add to system [0.0]

This simple script automates the process of generating a dummy offset file from <FITSFile> (with noise of ±<noise>), loading it into fpcal with the FITSFile, generating an Astrometric Model and viewing the results in the residuals.tcl utility.

If you have a FITS file with an appropriate binary table and fields (see § 3.2.1) and you want to learn to use the system, this script is a good place to start.

# Comm 4 Approach.

This is the approach used by TJF using OzPoz comm4.   This is for processing the GIRAFFE Raster data.   This approach analyses the data and collects information required for any future off line analysis.

## Input Data

We need the raster flux files produced by the GIRAFFE DRS and the first and last FITS files from the raster.  Only one FITS file anywhere in the raster is needed for the calibration software - but the first and last allow us to get the start and end time of the sequence for comparison to the configuration time (ACTMJD)..

### Extracting information from Headers.

You need to view the FITS header information for the two fits files.  For example, you could use the RTD to view the header. (“command view->FitsHeader”).  Note that there are multiple FITS headers in the file, one for each binary table.  We are interested in the primary header (HDU 1) and the OzPoz_table header.

Extract the values of the  PLATE, ACTMJD, ACTUTC and GIRAWLEN keywords from OzPoz binary header (ACTUTC is not available in earlier files).  Record these values.

If ACTUTC is not available, the equivalent value (ACTMJD as a string) may be obtainable from the FLAMES log files.  Search /vltdata/tmp/logFile and /vltdata/tmp/logFile*.bck for the ACTMJD (about 6 significant figures).  Often a line containing this value will include the time as a string.  This only works if the log files still exist.

Find the value of the DATE-OBS keyword from the "HDU 1" FITS header block for the first and last file to get the time of the first and last exposures

### First and last flux files.

The raster template takes three observations at the central position (first observation, middle observation and last observation).  Hence there are three flux files available for the central position.  This software will complain about this so remove the first and last raster flux files

## Determining best offset from fluxes.

To generate the offsets to the position in the raster with the maximum flux (for each fibre) you use the fluxOffsets command.   You will want to specify a cut to ensure that any fibre who’s maximum flux level is too low is removed.   The following command does this for the raster files specified and writes the output to the file offsets.txt.

fluxOffsets -gdhs -o offsets.txt -cut 50000    rast*.txt

Some experimenting with the cut value may be required (150000 was more appropriate for a tyco field.

## Basic Analysis

First, I often link one of the two  FITS file to a an simpler name.  Eg.

ln -s FLAMES_GIRAF_RAST292_0154.fits  data.fits

This allow us to use the name data.fits instead of the typically long original file name.

Now run fpcal

Load in the data into fpcal and check the model parameters used during fibre placement are what was expected.

load data.fits

offsets offsets.txt

mdlpars show

You may want to look at a plot of the raw data at this point. You will first need to do a fit (but not fitting anything) so that the relevant data is available.

fit

plot

Note that in this plot, the residuals display gives you the raw offsets whilst the offsets display gives you the offsets after the mean is removed.

Now get a fit of the origin (we are looking for patterns with this, but as yet they are unclear)

set fit origin

fit

Save the RMS and the values parameters 0 and 3 from result.

Now remove the mean offset and get RMS of the original fit.

set fit noall

remmeamoff

fit

Save  the resultant RMS.  It may be worth looking at the fit with “plot” command.

Now fit rotation and scale.  These are the most significant parameters.

set fit rotation

set fit scale

fit

Save the RMS and results for parameters 6 (rotation) and 7 (scale).

Remove offsets with large residuals (see “show residuals”) and fit again

clear resabove 150

fit

Again save RMS and results for parameters 6 and 7.  If of interest, plot the results.  These results are the values of interest – but note the variances on them given by the FIT – there is no point in changing parameters if the variance is above the size of the change..

## Distortion analysis.

The first parameter of the distortion model is effectively the plate scale – through the scale is also somewhat related to the other distortion parameters.  The basic sequence for modelling the non-wavelength dependent distortion parameters is

load data.fits

offsets offsets.txt

set fit rotation

set fit dist.

fit

But in rasters so far, the variances on the higher order distortion terms are very high.  A Powell fit may give better results (Use the command “set fitfun powell”) but it does not show variances and I am unconvinced.

### Wavelength distortion terms.

There are two wavelength dependent distortion terms.  Whilst the software attempts to provide facilities to analyse these, the variance on these terms is so high that it is hard to gain much improvement.  Hopefully the theoretical values will prove sufficient.

## Putting new values into the Robot.

### Raw fit results into Robot.

Each fit will produce a file named “OzPozModel<n>.sds where <n> is the plate number.  The result can be loaded into the robot software the wflames machine, account “flames” or “flamemgr”.

fpcolMdlLoad <n>

If you want to load the same model on both plates, then copy the file to the name for the other plate and execute the command for that plate.

### Hand edited results.

You may wish to combine the results of various analysis – averaging  results etc.  Say for example,you wish to insert a rotation of 0.1 degrees and a scale of 0.9996.  The calibration software can help with this.  For example

load data.fits

mdlpars show

show pars

mdlpars set 6 0.1

mdlpars set 7 0.9996

show pars

mdlpars show

mdlpars savesds

The “mdlpars show” command just gives you an analysis of the existing parameters.  “show pars” gives the existing parameter values, names and units.  It is best to add rotation and scale using the “extra rotation” and “extra scale” parmeters (numbered 6 and 7).  “mdlpars set” allows you to set a value of a parameter.  “mdlpars savesds” will normalize the model (putting the rotation in the the linear model parameters and the scale into the distortion model parameters).   The “mdlpars savesds” will generate the OzPozModel<n>.sds file which can now be used as per the previous section.

### Passing the new model to configure.

Configure (FPOSS) uses the telescope model and you should normally ensure it has access to the model the robot is using.  First you may need to use the template FP_NVRAM/FP_gen_tec_NVRAM_write in order to force the FP instrument status file to be written to disk.  This contains, amongst other things, the telescope model parameters.

Then, you need to copy the file /vltdata/config/fpInsStatus.sds from WFLAMES to the file of the same name in the directory you are running configure from or which contains the configure executable.   (Configure first looks in the current working directory for this file and then, if not found, in the executable directory).

## Software location and building.

The calibration software is archived in the fpcol CMM module.  It does not need the VLT CCS software to run but you will need it to extract the software from CMM.  Use “cmmCopy fpcol” to fetch the software.  Note, on the WFLAMES machine, this module is normally found in ~flamemgr/FLAMES/FP/ICS/fpcol.

The fpcol Module contains software common to the FP system, Configure software and calibration software.  To extract the calibration software, goto the fpcol/ws/src directory and execute

make cal

This creates a tar file named fpcolCal.tar.gz.  This can be moved else to build the software on any Unix machine with the SLA and CFITSIO libaries.  (Tested on HP-UX, Solaris and MacOS X).   You may have to edit the makefile, which is set up for HP-UX in a VLT configuration, to set the location of SLA and CFITSIO.  Then just execute “make”.

Typically, the software can be found in the directory ~flamemgr/FLAMES/FP/ICS/fpcol/ws/src/fpcolCal  on the WFLAMES machine.

## Details required for later analysis

If you wish to save the data for later analysis or pass it to somebody else via e-mail, you do not need to send the quite large FITS file. Instead, use the “savehd” command

___oOo___

The following C code shows how to read/write a row of the offset data:

scanf("%d%lf%lf", &fibre, &oRA, &oDEC);

printf("%d %18.15f %18.15f\n", fibre, oRA, oDEC);